## Slide 1: THE TRABECULAR METAL MONOBLOC ACETABULAR CUP DEMONSTRATES NO EARLY BEDDING-IN AND EXCELLENT CLINICAL PERFORMANCE AT 10-15 YEAR FOLLOW UP
ANTHONY S UNGER, MD
RANDALL J LEWIS, MD
JOHN MARTELL, MD

## Slide 2: ANTHONY S UNGER, MD
DISCLOSURES
  INNOMED—ROYALTIES
  BIOMET– ROYALTIES,CONSULTANT
  STRYKER—CONSULTANT, SPEAKER
  CDD-STOCK HOLDER, ROYALTIES
  J OF ARTHROPLASTY-EDITORIAL BOARD

## Slide 3
Trabecular Metal Acetabular Cup
Used Clinically Since 1997

## Slide 4
A pure metallic element, # 73.
A strong, ductile metal similar to titanium.
Tantalum is the most biocompatible metal of all.
50+ years of clinical use; cranioplasty plates, radio opaque markers, pace maker leads, ligation clips, etc...
TANTALUM

## Slide 5
Trabecular Metal
Properties
80 – 85% porosity
3-D open structure similar to bone
Uniform, repeating cellular structure
Consistent mechanical properties

## Slide 6
Trabecular Metal
Properties
Large unrestricted pathways for tissue ingrowth
Stiffness similar to bone
‘Stand-alone’; no solid metal needed

## Slide 7
Stiffness of Orthopaedic Materials
15
1.5
205
110
3
0.8
1.0
2.0
0
20
40
60
80
100
120
140
160
180
200
Cancellous Bone
Polyethylene
Subchondral Bone
Bone Cement
Trabecular Metal
Cortical Bone
Titanium Alloy
Co-Cr Alloy
Stiffness (GPa)
Trabecular Metal
Dense Cancellous Bone
Metals are 30 to 50X stiffer than bone & TM

## Slide 8
The Elliptical Profile
2 mm on the diameter
No bone displaced at dome
Spherical Preparation
Cup Elliptical Profile
Cup Elliptical Profile

## Slide 9
DCM Poly – Low Wear
Linear Wear mm/year
‘Problem Poly’ from literature
‘Gold Standard’
Poly from literature
Direct Molded Poly

## Slide 10
Direct Compression Molded Polyethylene
INTEGRATED
Into Trabecular Metal

## Slide 11
Direct Compression Molding of Poly Into the TMT Shell
1.5 to 2.0 mm of poly penetration into Trabecular Metal
Calcium stearate free UHMWPE resin…
…Is directly molded into the Trabecular Metal Shell
Trabecular Metal
Polyethylene
The ‘Interface’

## Slide 12
Backside poly wear & metal debris from locking mechanism
A Potential Source of Osteolysis
abrasion and screw holes with direct access to bone

## Slide 13
Pre-Op
CASE EXAMPLE
Rochester, MN
56 yr
Female
DJD

## Slide 14
Patient is pain-free with excellent function & ROM
2 years
Rochester, MN

## Slide 15: METHODS
347 PRIMARY THA IN 321 PTS
TWENTY BILATERAL THA
ALL SEEN AT MIN 10 YR FU
XRAYS AND CLINICAL EVALUATION AT 6 WKS, 6 MONTHS, YEARLY INTERVALS
HARRIS HIP SCORES

## Slide 16: METHODS
ANALOG IMAGES CONVERTED BY VIDAR DIAGNOSTIC PRO
DIGITAL IMAGES ANALYZED IN DICOM FORMAT
WEAR RATE ANLYSIS BY EDGE DETECTION METHOD OF MARTELL

## Slide 17: METHODS
BEDDING-IN DETERMINED BY WEAR RATE COMPARISON BETWEEN 6 WEEK TO ONE YEAR VERSUS ONE YEAR TO LONGEST FU
BEDDING-IN CONSIDERED NOT PRESENT WHEN NO SIGNIFICANT DIFFERENCE SEEN EARLY(ONE YEAR) WEAR AND LONGEST FU WEAR

## Slide 18: RESULTS
FU RANGE 10-15.4 YRS, MEAN 12.4
347 HIPS PERFORMED FROM 1997-2003
101 EXCLUDED DUE TO DEATH
108 LOST TO FU, OR TOO OLD TO RETURN
132 HIPS AVAILABLE FOR ANALYSIS

## Slide 19: RESULTS
HARRIS HIP SCORE 91 AT FU
NO CUP ASEPTIC LOOSENING
NO OSTEOLYSIS
NO REVISION FOR CUP LOOSENING
SIX HIPS REVISED FOR RECURRENT DISLOCATION

## Slide 20: WEAR ANALYSIS
ANTEVERSION MADE SOME CUPS IMPOSSIBLE TO VISUALIZE THE HEAD
WEAR ANALYSIS FOR BEDDING-IN PERFORMED IN 67 HIPS
TRUE WEAR ANALYSIS IN 55 HIPS

## Slide 21: WEAR ANALYSIS
 | BEDDING IN | TRUE WEAR | ONE WAY ANOVA | MANN-WHITNEY
#HIPS | 67 | 55 |  | 
MEAN LINEAR WEAR RATE(MM/YR) | 0.096 | 0.133 | P=0.44 | 
MEDIAN LINEAR WEAR RATE(MM/YR) | 0.099 | 0.099 |  | P=0.66
LINEAR STD DEV | .216 | 0.31 |  | 
MEAN VOL RATE(MM3/YR) | 55.5 | 84.26 | P=0.27 | 
MEDIAN VOL RATE(MM3/YR) | 22.1 | 18.01 |  | P=0.91
VOL STD DEV | 112 | 173 |  | 

## Slide 22: CONCLUSIONS
OVERALL WEAR RATE LOW
NO OSTEOLYSIS OR LOOSENING AT 15 YR FU
CLINICAL RESULTS EXCELLENT
BEDDING-IN DOES NOT OCCUR IN MONOBLOC DESIGNED CUP WITH DCM POLY